## Slide 1
Modern Myoelectric Intelligent Hand Prostheses
Tobias Stocker (Advisor: Pascal Weiner)
Seminar: Humanoid Robotics, WS 2017/18

## Slide 2: Outline
Motivation and Challenges
Overview of Hand Prostheses
Comparison of Hand Prostheses
Conclusion

## Slide 3: Motivation
Why myoelectric hand prostheses?
  Make a normal life possible for amputees
  Enable users to perform different grasps
  for activities of daily living (ADLs)
  Allow the user to control the hand through muscle contraction (with EMG)
  Electric actuators are rather small, quiet and
  have good precision and controllability

## Slide 4: Motivation
Desired properties:
  comfortable (lightweight, small)
  many different grasps possible
  High finger forces / fast joint speed
  easy to use
  high durability (robust, easy to repair)
  low-cost
  intelligent functions (sensor-feedback, grasp adaption)
Desired properties are contradicting trade-offs are mandatory

## Slide 5: Challenges
„Even state-of-the art devices lack a combination of high functionality, durability, adequate cosmetic appearance, and affordability“
Joseph T. Belter, 2013
Total weight should be below (human hand: )
lightweight materials, small and low number of actuators
transmission systems that allow for many different grasps
Finger tip force in precision grasp should be (human hand: )
and joint speed should be (human hand: )
Finger kinematic designs should be simple and robust
User should be able to move the hand without concentrating
User should get sensor feedback from the hand

## Slide 6: Transmission systems
Gears/belts
Four-bar linkage
Tendons

## Slide 7: Outline
Motivation and Challenges
Overview of Hand Prostheses
Comparison of Hand Prostheses
Conclusion

## Slide 8: Tact
Year | Mass (g) | Joints | DoF | Actuators | Transmission
2015 | 350 | 11 | 6 | 6 | Four-bar linkage
Special feature:
  Matches performance of other myoelectric prosthetic hands,
  while being very cheap ()
  Easy to manufacture with 3D-printer
  and of-the shelf parts
  Open-source

## Slide 9: Hand of Bennett et al.
Special feature:
  Four motor units in unique configuration:
    2 for thumb and 1 for index (fully actuated)
    1 for other fingers (underactuated)
  Embedded control system that enables
  self-contained control of hand movement
Year | Mass (g) | Joints | DoF | Actuators | Transmission
2015 | 546 | 12 | 12 | 4 | Tendons

## Slide 10: Hand of Bennett et al.
Embedded system:
  Consists of a single , four-layer circuit board
  Accepts and executes motion/force commands from a high-level controller
  Returns processed position/force information

## Slide 11: Hand of Zhang et al.
Special feature:
  Fingers are equipped with numerous
  torque and position sensors
  Integrated motion control system consisting of a motion control subsystem and several sensory subsystems
  New concept for sensory feedback system
  based on an electrical stimulator
Year | Mass (g) | Joints | DoF | Actuators | Transmission
2015 | 420 | 15 | 5 | 5 | Four-bar linkage

## Slide 12: Hand of Zhang et al.
Sensor system:
  Equipped with 18 proprioceptive and exteroceptive sensors
Embedded system:
  Consists of a motion control subsystem and several sensory subsystems

## Slide 13: SSSA-MyHand
Special feature:
  Only three actuators
  Abduction/adduction of the thumb and flexion/extension of the index with single actuator via Geneva drive
  Embedded controller and sensory system with force/position sensors and automatic grasp control
Year | Mass (g) | Joints | DoF | Actuators | Transmission
2016 | 478 | 10 | 4 | 3 | Four-bar linkage, Geneva drive

## Slide 14: SSSA-MyHand
Embedded system:
  Master-slave configuration based on a pair of 8-bit microcontrollers
  Supports:
    Identification of external commands
    Implementation of automatic motor functions
    Real time processing of the internal sensors
    Potential delivery of sensory feedback

## Slide 15: AstoHand v.1
Special feature:
  Low-cost
  Very lightweight
  Built with 3D-printed material
  (easy to manufacture and maintain)
Embedded system:
  Arduino Nano microcontroller
  Control algorithm developed in Simulink
  Digital input is used for selecting one of seven grip patterns
Year | Mass (g) | Joints | DoF | Actuators | Transmission
2016 | 261 | 10 | 5 | 5 | Tendons

## Slide 16: X-Hand
Special feature:
  Anthropomorphic grasping ability via special motion distribution mechanism structure
  Can replicate almost all natural movement of the human hand while using few actuators
Year | Mass (g) | Joints | DoF | Actuators | Transmission
2016 | - | 16 | 4 | 4 | Tendons

## Slide 17: Six-DOF-Hand
Special feature:
  Inexpensive
  Open source
  Independent finger movements
  Actuators with encoders for motor position feedback
Year | Mass (g) | Joints | DoF | Actuators | Transmission
2016 | 584 | 10 | 6 | 6 | Gears/Belts

## Slide 18: Bionic Hand
Special feature:
  Hybrid actuated with Brushless DC motors and Shape Memory Alloy (SMA)
  Close replication of the human hand (with all structures, joints and tendons)
  24 degrees of freedom
Year | Mass (g) | Joints | DoF | Actuators | Transmission
2016 | - | 24 | 24 | 13 | Tendons

## Slide 19: SoftHand Pro-D
Special feature:
  Strongly underactuated softhand
  19 joints with only one single actuator
  Can move along two different synergistic directions to perform either precision or power grasp
  Decoding of movement intensions using the dynamic frequency content
Year | Mass (g) | Joints | DoF | Actuators | Transmission
2016 | - | 19 | 19 | 1 | Tendons

## Slide 20: UOMPro
Special feature:
  Simple serial communication interface to link with high-level control methods
  The implemented low-level controller can handle individual finger position commands or hand grip pattern commands
Year | Mass (g) | Joints | DoF | Actuators | Transmission
2017 | 432 | 10 | 6 | 6 | Four-bar linkage

## Slide 21: UOMPro
Embedded system:
  Low-level controller for close loop control of the motors using various control techniques
  High-level controller to identify the motion intention

## Slide 22: MORA Hap-2
Special feature:
  Self-adaption ability
  Finger mechanism is capable of generating passively different flexion/extension angles
  Fingers have under-actuation mechanism
Year | Mass (g) | Joints | DoF | Actuators | Transmission
2017 | 250 | 14 | 11 | 4 | Four-bar linkage

## Slide 23: Outline
Motivation and Challenges
Overview of Hand Prostheses
Comparison of Hand Prostheses
Conclusion

## Slide 24: Physical properties
Name | Weight (g) | Size (length x width x thickness in mm)
Tact | 350 | 200 x 98 x 27
Hand of Bennett et al. | 546 | 200 x 89 x -
Hand of Zhang et al. | 420 | 159 x 79 x 21
MyHand | 478 | 200 x 84 x 56
AstoHand v.1 | 261 | 180 x 85 x 50
X-Hand | - | Human hand size
Six-DOF-Hand | 584 | 202 x 99 x 61
Bionic Hand | - | -
SoftHand Pro-D | - | -
UOMPro | 432 | 189 x 88 x -
MORA Hap-2 | 250 | 95 (fingers) x 83 x 25

## Slide 25: Finger Kinematics
Name | Number of Joints | Joints per Finger |  | 
 |  | Thumb | Index | Others
Tact | 11 | 3 | 2 | 2
Hand of Bennett et al. | 12 | 3 | 3 | 2
Hand of Zhang et al. | 15 | 3 | 3 | 3
MyHand | 10 | 2 | 2 | 2
AstoHand v.1 | 10 | 2 | 2 | 2
X-Hand | 16 | 4 | 3 | 3
Six-DOF-Hand | 10 | 2 | 2 | 2
Bionic Hand | 24 | 3 | 4 | 4
SoftHand Pro-D | 19 | 3 | 4 | 4
UOMPro | 10 | 2 | 2 | 2
MORA Hap-2 | 14 | 2 | 3 | 3

## Slide 26: Actuation and Transmission
Name | Actuators | Transmission system
Tact | 6 | Four-bar linkage
Hand of Bennett et al. | 4 | Tendons
Hand of Zhang et al. | 5 | Four-bar linkage
MyHand | 3 | Four-bar linkage, Geneva drive
AstoHand v.1 | 5 | Tendons
X-Hand | 4 | Tendons
Six-DOF-Hand | 6 | Gears / Belts
Bionic Hand | 13 | Tendons
SoftHand Pro-D | 1 | Tendons
UOMPro | 6 | Four-bar linkage
MORA Hap-2 | 4 | Four-bar linkage

## Slide 27: Grasping
Most of the hands are capable of performing the known grasping patterns:
  Power grasp (Cylindrical grasp)
  Precision grasp (Pinch grasp)
  Lateral grasp
  Hook grasp
The MORA Hap-2 is mainly developed for power and hook grasp
The SoftHand Pro-D with its single actuator can perform power and precision grasps

## Slide 28: Dynamics
Name | Individual Finger Force | Joint Speed
Tact | 4.2 N | 250 °/s
Hand of Bennett et al. | 25-30 N | -
Hand of Zhang et al. | 4.3-10 N | 68-118 °/s
MyHand | 12-31 N | 160-250 °/s
AstoHand v.1 | - | -
X-Hand | - | -
Six-DOF-Hand | 4.1 N | 128 °/s
Bionic Hand |  | -
SoftHand Pro-D | - | -
UOMPro | - | -
MORA Hap-2 | - | -

## Slide 29: Outline
Motivation and Challenges
Overview of Hand Prostheses
Comparison of Hand Prostheses
Conclusion

## Slide 30: Conclusion
There are still a lot of open problems in designing myoelectric hand prostheses
Many research groups develop prosthetic hands with new features
The main goal in the last years was to develop preferably low-cost protheses
Most research groups currently focus mainly on one novel design approach in their hand design
Only a few groups tried to incorporate intelligent functions

## Slide 31
Thank you for your attention!